Exam workflow:

- WF exam acquisition – 3 exams captured using new relaxed settings and blink detection. The software will give the user 5 capture attempts total and then stop the process, presenting the captured exams using the manual acceptance method below (see step 3).

- WF exam acceptance – SW will attempt to auto-accept the exam by comparing the 3 captures. SW Reviews the average Total RMS calculated for each and if each exam is within 1.00µ, the SW selects the exam capture with the median total RMS.

- If the exams vary greater than 1.00µ, the SW will present the points displays for each exam capture and allow the user to select their favorite.  Will need some kind of navigation built in to toggle between the 3 points displays and some kind of GUI to select the favorite.  If you have ideas here, we’re open to them.

- CT exam acquisition – if user selects the dual exam process, the CT exam acquisition will commence as soon as the 3 WF exams are captured and the WF acceptance will take place after CT and/or color image is acquired.  Otherwise, CT exam acquisition begins with user clicking the CT exam button.

- The new Ring coverage analysis occurs for the auto-acceptance.

- Color image is acquired (if selected in Settings).

- User is presented with the Acquisition Results display as described here.

E Line 1: Time date stamp in smaller font, OD(OS) label in larger font

D

C

B

A

Line 1: Time date stamp in smaller font, OD(OS) label in larger font

Acquisition Results Display, Dual Exam:

Panel A: WF eye image with entry points pattern.

Change the colors for the entry points to Green for good points and Red for rejected points.  Include both sets of 128 points on one eye if possible.  The dots will likely overlap.

Panel B: CT eye image with rings shown.

Panel C: WF Exam Acquisition Details:

Line 1: Time date stamp in smaller font, OD(OS) label in larger font. Continue to flag red if manually edited OD(OS).

Line2: Indicate whether Auto-Acceptance by SW or Manual-Acceptance by user, and which WF exam capture is used and how many total were captured and analyzed.

Line3: Centration method

Line4: Pupil / Scan size.  Indicate if the pupil size measured was larger than 6.00mm by flagging it with red font.

Line 5: Tracey Refraction measure with color coding as in past versions.

Panel D: CT Exam Acquisition Details:

Line 1: Time date stamp in smaller font, OD(OS) label in larger font. Continue to flag red if manually edited OD(OS).

Line2: Indicate the result of the ring analysis (I’m not familiar with how this is implemented so I’m not sure what to include here. If there is an auto acceptance method vs a manual acceptance, it should indicate here which one was used.)

Line3: Ks as selected in Settings (WF, Refractive or SimK). We’ll need to add WF Ks as an option to display where Ks are displayed – I need to create another requirement for this).

Panel E: New Eye icon

| If | And if | Display |
| --- | --- | --- |
| Cornea HO <0.100 µ | Internal HO <0.100 µ | CG-IG.png |
| Cornea HO <0.100 µ | Internal HO >=0.100µ and <0.200µ | CG-IY.png |
| Cornea HO <0.100 µ | Internal HO >=0.200µ | CG-IR.png |
| Cornea HO >=0.100µ and <0.200µ | Internal HO <0.100 µ | CY-IG.png |
| Cornea HO >=0.200µ | Internal HO <0.100 µ | CR-IG.png |
| Cornea HO >=0.100µ and <0.200µ | Internal HO >=0.100µ and <0.200µ | CY-IY.png |
| Cornea HO >=0.100µ and <0.200µ | Internal HO >=0.200µ | CY-IR.png |
| Cornea HO >=0.200µ | Internal HO >=0.100µ and <0.200µ | CR-IY.png |
| Cornea HO >=0.200µ | Internal HO >=0.200µ | CR-IR.png |

The Navigation bar along the right should have only one display option as Summary which will go to the new Summary display.

Acquisition Results Display, WF Exam:

E Line 1: Time date stamp in smaller font, OD(OS) label in larger font

A

B

C

D

Panel A: WF eye image with entry points pattern.

Change the colors for the entry points to Green for good points and Red for rejected points.  Include both sets of 128 points on one eye if possible.  The dots will likely overlap.

Panel B: NoExam.png displayed in the panel to indicate there is no CT eye exam to display.

Panel C: WF Exam Acquisition Details:

Line 1: Time date stamp in smaller font, OD(OS) label in larger font. Continue to flag red if manually edited OD(OS).

Line2: Indicate whether Auto-Acceptance by SW or Manual-Acceptance by user, and which WF exam capture is used and how many total were captured and analyzed.

Line3: Centration method

Line4: Pupil / Scan size.  Indicate if the pupil size measured was larger than 6.00mm by flagging it with red font.

Line 5: Tracey Refraction measure with color coding as in past versions.

Panel D: CT Exam Not Acquired with this WF

Panel E: New Eye icon

Display the grayed out eye icon – EyeGrey.png

The Navigation bar along the right should have only one display option as Summary which will go to the new Summary display.

Acquisition Results Display, CT Exam:

E Line 1: Time date stamp in smaller font, OD(OS) label in larger font

A

D

B

C

Panel A: NoExam.png displayed in the panel to indicate there is no WF eye exam to display.

Panel B: CT eye image with rings shown.

Panel C: WF Exam Not Acquired with this CT:

Panel D: CT Exam Acquisition Details:

Line 1: Time date stamp in smaller font, OD(OS) label in larger font. Continue to flag red if manually edited OD(OS).

Line2: Indicate the result of the ring analysis (I’m not familiar with how this is implemented so I’m not sure what to include here. If there is an auto acceptance method vs a manual acceptance, it should indicate here which one was used.)

Line3: Ks as selected in Settings (WF, Refractive or SimK). We’ll need to add WF Ks as an option to display where Ks are displayed – I need to create another requirement for this).

Panel E: New Eye icon

| If | Display |
| --- | --- |
| Cornea HO <0.100 µ | CG.png |
| Cornea HO >=0.100µ and <0.200µ | CY.png |
| Cornea HO >=0.200µ | CR.png |

The Navigation bar along the right should have only one display option as Summary which will go to the new Summary display.